## Slide 1
Wafermap Analysis Report
Generated on 2026-02-26 14:24

## Slide 2
Gage R&R Report - Agenda
# | Group | Parameters | Pages | Slides
1 | 📋 Agenda | - | Page 2 | 1
2 | 📖 Calculation Parameters | - | Page 3 | 1
3 | 📍 Stitched - Pixel Position based | 5 dies | Page 4-9 | 6
Total: 3 Groups | 0 Parameters | 8 Slides

## Slide 3
Gage R&R - Calculation Parameters
BASIC STATISTICAL VALUES
Mean (Mittelwert):
Average of all measurements across all files and dies.
Mean = Σ(xi) / n
Std Dev (Standardabweichung):
Measure of dispersion around the mean.
σ = √[Σ(xi - Mean)² / n]
Range (Spannweite):
Difference between maximum and minimum.
Range = Max - Min
LSL / USL (Spec Limits):
Statistical limits (k = sigma factor).
LSL = Mean - kσ | USL = Mean + kσ
GAGE R&R COMPONENTS
Repeatability (EV):
Variation within measurements of one file (equipment).
EV = Mean(σ within each file)
Reproducibility (AV):
Variation between files/operators/trials.
AV = σ(file means)
GRR Total:
Combined measurement system variation.
GRR = R̄ / d₂ (Range Method)
ASSESSMENT METRICS
%GRR (Prozent GRR):
GRR as percentage of total variation or tolerance.
%GRR = (GRR / TV) × 100 or (GRR / Tolerance) × 100
ndc (Using Total Variation):
Number of distinct categories the system can distinguish.
ndc = 1.41 × (TV / GRR)
ASSESSMENT CRITERIA (AIAG MSA)
🟢 EXCELLENT: %GRR < 10%
Measurement system is acceptable.
🟡 ACCEPTABLE: 10% ≤ %GRR < 30%
May be acceptable depending on application.
🔴 UNACCEPTABLE: %GRR ≥ 30%
Measurement system must be improved.
ndc ≥ 5: Good measurement system
Can distinguish ≥ 5 distinct categories.
INTERPRETATION
• High Repeatability → Equipment/method problem
• High Reproducibility → Operator/process/time problem
• ndc should be ≥ 5 for a good measurement system
• Values outside LSL/USL indicate outliers

## Slide 4
PLM Pixel Analysis - Stitched - Die (33,33)
Region | Position | Size | %GRR | ndc | Repeat.% | Reprod.% | Result
R1 | (96,94)->(121,119) | 25x25 | 24.0% | 5.7 | 23.7% | 4.1% | ACCEPTABLE
R2 | (445,131)->(470,156) | 25x25 | 23.6% | 5.8 | 23.4% | 3.4% | ACCEPTABLE
R3 | (264,267)->(289,292) | 25x25 | 28.2% | 4.8 | 28.1% | 2.8% | ACCEPTABLE
R4 | (131,431)->(156,456) | 25x25 | 26.2% | 5.2 | 25.2% | 7.1% | ACCEPTABLE
R5 | (313,488)->(338,513) | 25x25 | 24.2% | 5.7 | 23.7% | 4.9% | ACCEPTABLE
R6 | (426,361)->(451,386) | 25x25 | 21.2% | 6.5 | 20.5% | 5.5% | ACCEPTABLE
R7 | (248,632)->(273,657) | 25x25 | 23.7% | 5.8 | 22.1% | 8.5% | ACCEPTABLE
Die (33, 33)
PLM Type: Stitched
Regions: 7
Files/Trials: 2
Pixels/Region: 625
Best: R6 - %GRR=21.2%
Worst: R3 - %GRR=28.2%

## Slide 5
PLM Pixel Analysis - Stitched - Die (35,44)
Region | Position | Size | %GRR | ndc | Repeat.% | Reprod.% | Result
R1 | (96,94)->(121,119) | 25x25 | 18.0% | 7.7 | 14.5% | 10.7% | ACCEPTABLE
R2 | (445,131)->(470,156) | 25x25 | 20.4% | 6.8 | 16.7% | 11.7% | ACCEPTABLE
R3 | (264,267)->(289,292) | 25x25 | 16.8% | 8.3 | 11.8% | 11.9% | ACCEPTABLE
R4 | (131,431)->(156,456) | 25x25 | 16.2% | 8.6 | 13.2% | 9.4% | ACCEPTABLE
R5 | (313,488)->(338,513) | 25x25 | 16.9% | 8.2 | 14.6% | 8.6% | ACCEPTABLE
R6 | (426,361)->(451,386) | 25x25 | 16.1% | 8.6 | 12.0% | 10.7% | ACCEPTABLE
R7 | (248,632)->(273,657) | 25x25 | 18.8% | 7.4 | 17.4% | 7.2% | ACCEPTABLE
Die (35, 44)
PLM Type: Stitched
Regions: 7
Files/Trials: 2
Pixels/Region: 625
Best: R6 - %GRR=16.1%
Worst: R2 - %GRR=20.4%

## Slide 6
PLM Pixel Analysis - Stitched - Die (44,44)
Region | Position | Size | %GRR | ndc | Repeat.% | Reprod.% | Result
R1 | (96,94)->(121,119) | 25x25 | 15.3% | 9.1 | 11.3% | 10.3% | ACCEPTABLE
R2 | (445,131)->(470,156) | 25x25 | 15.7% | 8.9 | 12.3% | 9.6% | ACCEPTABLE
R3 | (264,267)->(289,292) | 25x25 | 15.0% | 9.3 | 10.1% | 11.1% | ACCEPTABLE
R4 | (131,431)->(156,456) | 25x25 | 14.7% | 9.5 | 12.3% | 8.1% | ACCEPTABLE
R5 | (313,488)->(338,513) | 25x25 | 13.0% | 10.8 | 8.6% | 9.8% | ACCEPTABLE
R6 | (426,361)->(451,386) | 25x25 | 12.3% | 11.4 | 8.6% | 8.8% | ACCEPTABLE
R7 | (248,632)->(273,657) | 25x25 | 15.1% | 9.3 | 12.4% | 8.5% | ACCEPTABLE
Die (44, 44)
PLM Type: Stitched
Regions: 7
Files/Trials: 2
Pixels/Region: 625
Best: R6 - %GRR=12.3%
Worst: R2 - %GRR=15.7%

## Slide 7
PLM Pixel Analysis - Stitched - Die (46,62)
Region | Position | Size | %GRR | ndc | Repeat.% | Reprod.% | Result
R1 | (96,94)->(121,119) | 25x25 | 21.2% | 6.5 | 17.2% | 12.3% | ACCEPTABLE
R2 | (445,131)->(470,156) | 25x25 | 21.1% | 6.5 | 16.3% | 13.3% | ACCEPTABLE
R3 | (264,267)->(289,292) | 25x25 | 22.0% | 6.3 | 19.6% | 10.0% | ACCEPTABLE
R4 | (131,431)->(156,456) | 25x25 | 19.0% | 7.3 | 16.5% | 9.5% | ACCEPTABLE
R5 | (313,488)->(338,513) | 25x25 | 21.0% | 6.6 | 16.1% | 13.4% | ACCEPTABLE
R6 | (426,361)->(451,386) | 25x25 | 22.5% | 6.1 | 19.0% | 12.1% | ACCEPTABLE
R7 | (248,632)->(273,657) | 25x25 | 20.2% | 6.8 | 17.7% | 9.7% | ACCEPTABLE
Die (46, 62)
PLM Type: Stitched
Regions: 7
Files/Trials: 2
Pixels/Region: 625
Best: R4 - %GRR=19.0%
Worst: R6 - %GRR=22.5%

## Slide 8
PLM Pixel Analysis - Stitched - Die (56,58)
Region | Position | Size | %GRR | ndc | Repeat.% | Reprod.% | Result
R1 | (96,94)->(121,119) | 25x25 | 12.6% | 11.1 | 9.0% | 8.8% | ACCEPTABLE
R2 | (445,131)->(470,156) | 25x25 | 13.8% | 10.1 | 10.6% | 8.8% | ACCEPTABLE
R3 | (264,267)->(289,292) | 25x25 | 16.5% | 8.4 | 16.2% | 3.2% | ACCEPTABLE
R4 | (131,431)->(156,456) | 25x25 | 14.6% | 9.6 | 13.9% | 4.4% | ACCEPTABLE
R5 | (313,488)->(338,513) | 25x25 | 15.7% | 8.8 | 14.4% | 6.4% | ACCEPTABLE
R6 | (426,361)->(451,386) | 25x25 | 15.4% | 9.1 | 13.9% | 6.5% | ACCEPTABLE
R7 | (248,632)->(273,657) | 25x25 | 14.3% | 9.7 | 13.9% | 3.5% | ACCEPTABLE
Die (56, 58)
PLM Type: Stitched
Regions: 7
Files/Trials: 2
Pixels/Region: 625
Best: R1 - %GRR=12.6%
Worst: R3 - %GRR=16.5%

## Slide 9
PLM Pixel Analysis - Stitched Overview - Stitched
Die | R1 %GRR | R2 %GRR | R3 %GRR | R4 %GRR | R5 %GRR | R6 %GRR | R7 %GRR | Avg %GRR
(33,33) | 24.0% | 23.6% | 28.2% | 26.2% | 24.2% | 21.2% | 23.7% | 24.4%
(35,44) | 18.0% | 20.4% | 16.8% | 16.2% | 16.9% | 16.1% | 18.8% | 17.6%
(44,44) | 15.3% | 15.7% | 15.0% | 14.7% | 13.0% | 12.3% | 15.1% | 14.4%
(46,62) | 21.2% | 21.1% | 22.0% | 19.0% | 21.0% | 22.5% | 20.2% | 21.0%
(56,58) | 12.6% | 13.8% | 16.5% | 14.6% | 15.7% | 15.4% | 14.3% | 14.7%
Stitched Die Comparison
PLM Type: Stitched
Total Dies: 5
Regions per Die: 7
Files/Trials: 2
Best: Die(44,44) R6 - 12.3%
Worst: Die(33,33) R3 - 28.2%